## example to follow
sample | sample mass | alkalinity | notes | mean | sd | sd*2 | cv | percent | min | max | range
alkcrm1 | 50.749 | 2242.6988 | all samples measured in order and results calculated at 25degC. Alkalinity CRM was from batch 130, an older batch donated from the Stillman lab. Goal was to get 2238.04 +- 0.53 umol/kg
alkcrm2 | 54.497 | 2247.5364
alkcrm3 | 53.735 | 2230.1036 |  | 2240.1129 | 8.9995 | 17.9989 | 0.0040174 | 0.4017% | 2230.1036 | 2247.5364 | 17.4327
BI-0075-P1-1 | 54.156 | 2114.1482
BI-0075-P1-2 | 52.368 | 2138.9194
BI-0075-P1-3 | 54.708 | 2142.2411 |  | 2131.7695 | 15.3507 | 30.7013 | 0.0072009 | 0.7201% | 2114.1482 | 2142.2411 | 28.0929
BI-0075-P2-1 | 53.209 | 2135.239
BI-0075-P2-2 | 53.708 | 2133.4876
BI-0075-P2-3 | 51.482 | 2129.4654 |  | 2132.7306 | 2.9603 | 5.9205 | 0.001388 | 0.1388% | 2129.4654 | 2135.239 | 5.7736
BI-0075-P3-1 | 50.766 | 2124.9361
BI-0075-P3-2 | 55.265 | 2132.3922
BI-0075-P3-3 | 51.326 | 2126.4715 |  | 2127.9333 | 3.9371 | 7.8742 | 0.0018502 | 0.185% | 2124.9361 | 2132.3922 | 7.4561
alkcrm4 | 55.583 | 2249.3508
alkcrm5 | 51.356 | 2250.2916
alkcrm6 | 53.685 | 2272.6196 | Sample sat overnight without being processed because the instrument arm couldn't settle correctly onto sample cup. | 2257.4207 | 13.1711 | 26.3421 | 0.0058346 | 0.5835% | 2249.3508 | 2272.6196 | 23.2688
|  |  |  | Cref value | CLabi-Cref | (xi-Cref)^2
|  |  |  | 2238.04 | 4.6588 | 21.7042
|  |  |  | 2238.04 | 9.4964 | 90.181
|  |  |  | 2238.04 | -7.9364 | 62.9861
|  |  |  | 2238.04 | 11.3108 | 127.9349
|  |  |  | 2238.04 | 12.2516 | 150.1007
[6 empty] | 9.5174 | RMSbias
[6 empty] | 0.53 | u(Cref)
[6 empty] | 9.5322 | u(bias) of x | absolute uncertainty
[6 empty] | 0.0042592 | u(bias)/x | relative uncertainty
https://www.c-can.info/reference/OAGuidanceManual.pdf
https://sisu.ut.ee/sites/default/files/measurement/files/single-lab_validation_nordtest_uncertainty.pdf

## data 11dec2019 to 17nov2020
sample | alkalinity at 25degC | sample mass | notes | mean | sd | sd*2 | cv | percent | min | max | range
crm130_12112019_1 | 2242.6988 | 50.749 | all samples measured in order and results calculated at 25degC. Alkalinity CRM was from batch 130, an older batch donated from the Stillman lab. Goal was to get 2238.04 +- 0.53 umol/kg
crm130_12112019_2 | 2247.5364 | 54.497
crm130_12112019_3 | 2230.1036 | 53.735 |  | 2240.1129 | 8.9995 | 17.9989 | 0.0040174 | 0.4017% | 2230.1036 | 2247.5364 | 17.4327
BI-0075-P1-1 | 2114.1482 | 54.156
BI-0075-P1-2 | 2138.9194 | 52.368
BI-0075-P1-3 | 2142.2411 | 54.708 |  | 2131.7695 | 15.3507 | 30.7013 | 0.0072009 | 0.7201% | 2114.1482 | 2142.2411 | 28.0929
BI-0075-P2-1 | 2135.239 | 53.209
BI-0075-P2-2 | 2133.4876 | 53.708
BI-0075-P2-3 | 2129.4654 | 51.482 |  | 2132.7306 | 2.9603 | 5.9205 | 0.001388 | 0.1388% | 2129.4654 | 2135.239 | 5.7736
BI-0075-P3-1 | 2124.9361 | 50.766
BI-0075-P3-2 | 2132.3922 | 55.265
BI-0075-P3-3 | 2126.4715 | 51.326 |  | 2127.9333 | 3.9371 | 7.8742 | 0.0018502 | 0.185% | 2124.9361 | 2132.3922 | 7.4561
crm130_12112019_4 | 2249.3508 | 55.583
crm130_12112019_5 | 2250.2916 | 51.356
crm130_12112019_6 | 2272.6196 | 53.685 | Sample sat overnight without being processed because the instrument arm couldn't settle correctly onto sample cup. | 2257.4207 | 13.1711 | 26.3421 | 0.0058346 | 0.5835% | 2249.3508 | 2272.6196 | 23.2688
crm162_01272020_1 | 2356.3091 |  | ran this cup first on autotitrator in cup slot 5, was labeled alkcrm2. renamed to new naming system which includes batch number and date ran. Acted as sacrificial sample so not great for CRM assessment
crm162_01272020_2 | 2403.7904 |  | alkcrm1 was ran afterwards in cup slot 6. renamed here to refelect new naming system | 2380.0497 | 33.5743 | 67.1487 | 0.014107 | 1.411% | 2356.3091 | 2403.7904 | 47.4813
R2T1-1 | 2089.0546
R2T1-2 | 2131.3932
R2T1-3 | 2117.0682 |  |  | 2112.5053 | 21.535 | 43.0699 | 0.010194 | 1.019% | 2089.0546 | 2131.3932 | 42.3386
R2T2-1 | 2120.9946
R2T2-2 | 2116.5703
R2T2-3 | 2119.6642 |  |  | 2119.0764 | 2.27 | 4.54 | 0.0010712 | 0.1071% | 2116.5703 | 2120.9946 | 4.4243
crm130_01272020_1 | 2278.2466
crm130_01272020_2 | 2283.2413 |  |  | 2280.7439 | 3.5318 | 7.0636 | 0.0015485 | 0.1549% | 2278.2466 | 2283.2413 | 4.9947
SACRIFICE | 1667.1035
crm130_01282020_1 | 2272.6035
crm186_01282020_1 | 2475.0735 |  |  | 2373.8385 | 143.1679 | 286.3359 | 0.060311 | 6.031% | 2272.6035 | 2475.0735 | 202.47
R3T2-1 | 2101.001
R3T2-2 | 2096.4751
R3T2-3 | 2092.9998 |  |  | 2096.8253 | 4.0121 | 8.0242 | 0.0019134 | 0.1913% | 2092.9998 | 2101.001 | 8.0012
R3T1-1 | 2094.7349
R3T1-2 | 2102.146
R3T1-3 | 2092.949 |  |  | 2096.61 | 4.8768 | 9.7536 | 0.002326 | 0.2326% | 2092.949 | 2102.146 | 9.197
crm186_01282020_2 | 2233.4604
crm186_01282020_3 | 2242.3757 |  |  | 2237.918 | 6.3041 | 12.6082 | 0.0028169 | 0.2817% | 2233.4604 | 2242.3757 | 8.9153
SACRIFICE | 1624.2912
crm186_01292020_1 | 2241.6865
crm186_01292020_2 | 2241.5421 |  |  | 2241.6143 | 0.10215 | 0.2043 | 4.5571e-05 | 0.004557% | 2241.5421 | 2241.6865 | 0.14446
R1T1-1 | 2125.8972
R1T1-2 | 2109.3232
R1T1-3 | 2115.631 |  |  | 2116.9504 | 8.3654 | 16.7308 | 0.0039516 | 0.3952% | 2109.3232 | 2125.8972 | 16.574
R1T2-1 | 2122.2192
R1T2-2 | 2123.5923
R1T2-3 | 2111.1737 |  |  | 2118.9951 | 6.8082 | 13.6164 | 0.0032129 | 0.3213% | 2111.1737 | 2123.5923 | 12.4186
crm186_01292020_3 | 2236.1096
crm186_01292020_4 | 2235.1851 |  |  | 2235.6473 | 0.65371 | 1.3074 | 0.00029241 | 0.02924% | 2235.1851 | 2236.1096 | 0.92449
SACRIFICE | 1592.3683
crm186_01302020_1 | 2248.0026
crm186_01202020_2 | 2247.3779 |  |  | 2247.6903 | 0.44178 | 0.88356 | 0.00019655 | 0.01965% | 2247.3779 | 2248.0026 | 0.62477
R1T3-1 | 2220.5347
R1T3-2 | 2220.7613
R1T3-3 | 2217.0855 |  |  | 2219.4605 | 2.06 | 4.1199 | 0.00092813 | 0.09281% | 2217.0855 | 2220.7613 | 3.6758
R1T4-1 | 2216.4759
R1T4-2 | 2217.8125
R1T4-3 | 2213.1885 |  |  | 2215.8256 | 2.3796 | 4.7592 | 0.0010739 | 0.1074% | 2213.1885 | 2217.8125 | 4.624
R2T3-1 | 2224.2949
R2T3-2 | 2223.8308
R2T3-3 | 2223.9694 |  |  | 2224.0317 | 0.23827 | 0.47654 | 0.00010713 | 0.01071% | 2223.8308 | 2224.2949 | 0.46417
crm186_01302020_3 | 2237.5032
crm186_01202020_4 | 2238.3866 |  |  | 2237.9449 | 0.62469 | 1.2494 | 0.00027914 | 0.02791% | 2237.5032 | 2238.3866 | 0.88344
SACRIFICE | 1592.1394
crm186_01312020_1 | 2247.372
crm186_01312020_2 | 2244.8768 |  |  | 2246.1244 | 1.7644 | 3.5288 | 0.00078554 | 0.07855% | 2244.8768 | 2247.372 | 2.4953
R3T3-1 | 2234.7436
R3T3-2 | 2227.2541
R3T3-3 | 2237.7694 |  |  | 2233.2557 | 5.4132 | 10.8265 | 0.0024239 | 0.2424% | 2227.2541 | 2237.7694 | 10.5153
R3T4-1 | 2233.5328
R3T4-2 | 2235.2718
R3T4-3 | 2229.9784 |  |  | 2232.9277 | 2.6981 | 5.3961 | 0.0012083 | 0.1208% | 2229.9784 | 2235.2718 | 5.2934
R2T4-1 | 2231.4122
R2T4-2 | 2218.8753
R2T4-3 | 2231.0744 |  |  | 2227.1207 | 7.1426 | 14.2853 | 0.0032071 | 0.3207% | 2218.8753 | 2231.4122 | 12.5369
crm186_01312020_3 | 2241.6594
crm186_01312020_4 | 2244.2021 |  |  | 2242.9308 | 1.798 | 3.596 | 0.00080163 | 0.08016% | 2241.6594 | 2244.2021 | 2.5428
SACRIFICE | 1604.1768
crm186_02102020_1 | 2169.8228
crm186_02102020_2 | 2189.8049 |  |  | 2179.8138 | 14.1295 | 28.2589 | 0.006482 | 0.6482% | 2169.8228 | 2189.8049 | 19.9821
bay001_02102020_1 | 2150.3665 |  | First baywater standard batch created
bay001_02102020_2 | 2141.5062 |  | First baywater standard batch created
bay001_02102020_3 | 2140.6435 |  | First baywater standard batch created | 2144.1721 | 5.3819 | 10.7637 | 0.00251 | 0.251% | 2140.6435 | 2150.3665 | 9.723
crm186_02102020_3 | 2207.0936
crm186_02102020_4 | 2203.9788 |  |  | 2205.5362 | 2.2025 | 4.4049 | 0.00099861 | 0.09986% | 2203.9788 | 2207.0936 | 3.1148
SACRIFICE | 1767.4128
bay001_02122020_1 | 1957.658
bay001_02122020_2 | 2158.5898 |  |  | 2058.1239 | 142.0802 | 284.1605 | 0.069034 | 6.903% | 1957.658 | 2158.5898 | 200.9318
crm186_02122020_1 | 2223.7492
crm186_02122020_2 | 2228.035 |  |  | 2225.8921 | 3.0305 | 6.061 | 0.0013615 | 0.1361% | 2223.7492 | 2228.035 | 4.2858
AMB-3-1 | 2176.5029
AMB-3-2 | 2186.0851
AMB-3-3 | 2212.4999 |  |  | 2191.696 | 18.6429 | 37.2858 | 0.0085061 | 0.8506% | 2176.5029 | 2212.4999 | 35.997
AMB-5-1 | 2103.7937
AMB-5-2 | 2076.5451
AMB-5-3 | 2102.5774 |  |  | 2094.3054 | 15.3929 | 30.7858 | 0.0073499 | 0.735% | 2076.5451 | 2103.7937 | 27.2486
bay001_02122020_3 | 2157.7459
bay001_02122020_4 | 2174.6168 |  |  | 2166.1814 | 11.9295 | 23.859 | 0.0055072 | 0.5507% | 2157.7459 | 2174.6168 | 16.8709
SACRIFICE | 1632.3655
bay001_02132020_1 | 2166.2382
bay001_02132020_2 | 2173.7654 |  |  | 2170.0018 | 5.3225 | 10.6451 | 0.0024528 | 0.2453% | 2166.2382 | 2173.7654 | 7.5272
crm186_02132020_1 | 2229.4135
ACID-3-1 | 2135.5996
ACID-3-2 | 2133.1924
ACID-3-3 | 2140.299 |  |  | 2136.3637 | 3.6144 | 7.2288 | 0.0016918 | 0.1692% | 2133.1924 | 2140.299 | 7.1066
ACID-5-1 | 2113.9334
ACID-5-2 | 2122.5769
ACID-5-3 | 2126.1611 |  |  | 2120.8905 | 6.2859 | 12.5718 | 0.0029638 | 0.2964% | 2113.9334 | 2126.1611 | 12.2277
bay001_02132020_3 | 2192.0776
bay001_02132020_4 | 2296.2887 |  |  | 2244.1831 | 73.6883 | 147.3767 | 0.032835 | 3.284% | 2192.0776 | 2296.2887 | 104.211
SACRIFICE | 1712.826
bay001_02172020_1 | 2156.3511
bay001_02172020_2 | 2167.2566
bay001_02172020_3 | 2161.0007 |  |  | 2161.5361 | 5.4724 | 10.9448 | 0.0025317 | 0.2532% | 2156.3511 | 2167.2566 | 10.9054
ACID-4-1 | 2095.2448
ACID-4-2 | 2100.3121
ACID-4-3 | 2089.5058 |  |  | 2095.0209 | 5.4066 | 10.8133 | 0.0025807 | 0.2581% | 2089.5058 | 2100.3121 | 10.8063
ACID-2-1 | 2218.8994
ACID-2-2 | 2186.6046
ACID-2-3 | 2189.3203 |  |  | 2198.2747 | 17.913 | 35.826 | 0.0081487 | 0.8149% | 2186.6046 | 2218.8994 | 32.2948
ACID-6-1 | 2186.4507
ACID-6-2 | 2194.8913
ACID-6-3 | 2193.0983 |  |  | 2191.4801 | 4.4469 | 8.8938 | 0.0020292 | 0.2029% | 2186.4507 | 2194.8913 | 8.4406
bay001_02172020_4 | 2173.0418
bay001_02172020_5 | 2182.9605
bay001_02172020_6 | 2181.0465 |  |  | 2179.0163 | 5.2618 | 10.5236 | 0.0024148 | 0.2415% | 2173.0418 | 2182.9605 | 9.9187
crm186_02172020_1 | 2232.7276
crm186_02172020_2 | 2225.3939
crm186_02172020_3 | 2224.1164 |  |  | 2227.4126 | 4.647 | 9.294 | 0.0020863 | 0.2086% | 2224.1164 | 2232.7276 | 8.6112
SACRIFICE | 1824.4795
BAYSTD1-02192020-1 | 2172.5331
BAYSTD1-02192020-2 | 2169.1763
BAYSTD1-02192020-3 | 2166.3611 |  |  | 2169.3568 | 3.09 | 6.1799 | 0.0014244 | 0.1424% | 2166.3611 | 2172.5331 | 6.172
AMB-1-1 | 2106.8453
AMB-1-2 | 2103.1035
AMB-1-3 | 2104.6341 |  |  | 2104.8609 | 1.8812 | 3.7624 | 0.00089375 | 0.08937% | 2103.1035 | 2106.8453 | 3.7419
ACID-1-1 | 2186.3282
ACID-1-2 | 2227.4413
ACID-1-3 | 2285.6798 |  |  | 2233.1498 | 49.9212 | 99.8424 | 0.022355 | 2.235% | 2186.3282 | 2285.6798 | 99.3516
BAYSTD1-02192020-4 | 2170.5506
BAYSTD1-02192020-5 | 2168.9045
BAYSTD1-02192020-6 | 2162.9156 |  |  | 2167.4569 | 4.0181 | 8.0362 | 0.0018538 | 0.1854% | 2162.9156 | 2170.5506 | 7.635
CRM186-02192020-1 | 2231.5014
CRM186-02192020-2 | 2237.2213
CRM186-02192020-3 | 2231.3102 |  |  | 2233.3443 | 3.3589 | 6.7179 | 0.001504 | 0.1504% | 2231.3102 | 2237.2213 | 5.9111
SACRIFICE | 1732.2672
BAYSTD1-02202020-1 | 2171.6037
BAYSTD1-02202020-2 | 2168.3325
BAYSTD1-02202020-3 | 2169.6807 |  |  | 2169.8723 | 1.644 | 3.288 | 0.00075765 | 0.07576% | 2168.3325 | 2171.6037 | 3.2712
AMB-6-1 | 2205.2934
AMB-6-2 | 2203.5716
AMB-6-3 | 2203.6333 |  |  | 2204.1661 | 0.97676 | 1.9535 | 0.00044314 | 0.04431% | 2203.5716 | 2205.2934 | 1.7218
AMB-4-1 | 2209.2654
AMB-4-2 | 2205.6469
AMB-4-3 | 2207.2129 |  |  | 2207.3751 | 1.8147 | 3.6294 | 0.00082211 | 0.08221% | 2205.6469 | 2209.2654 | 3.6185
AMB-2-1 | 2201.9531
AMB-2-2 | 2201.7544
AMB-2-3 | 2206.7819 |  |  | 2203.4965 | 2.847 | 5.694 | 0.001292 | 0.1292% | 2201.7544 | 2206.7819 | 5.0275
BAYSTD1-02202020-4 | 2172.3969
BAYSTD1-02202020-5 | 2167.4397
BAYSTD1-02202020-6 | 2171.8156 |  |  | 2170.5507 | 2.7099 | 5.4197 | 0.0012485 | 0.1248% | 2167.4397 | 2172.3969 | 4.9572
CRM186-02202020-1 | 2234.9085
CRM186-02202020-2 | 2233.8319
CRM186-02202020-3 | 2236.8026 |  |  | 2235.181 | 1.504 | 3.008 | 0.00067288 | 0.06729% | 2233.8319 | 2236.8026 | 2.9707
SAC-1 | 2791.4745
P-0078-1 | 2253.5401
P-0078-2 | 2250.349
P-0078-3 | 2261.7392
P-0078-4 | 2241.4624
P-0078-5 | 2258.1305
P-0078-6 | 2251.4238
P-0078-7 | 2254.6025 |  |  | 2253.0354 | 6.4314 | 12.8627 | 0.0028545 | 0.2855% | 2241.4624 | 2261.7392 | 20.2768
SAC-2 | 2424.3536
CRM186-11102020-1 | 2220.1115
CRM186-11102020-2 | 2223.2986
CRM186-11102020-3 | 2217.954 |  |  | 2220.4547 | 2.6888 | 5.3776 | 0.0012109 | 0.1211% | 2217.954 | 2223.2986 | 5.3447
SAC-1 | 2053.3156
P-0023-1 | 2121.1126
P-0023-2 | 2114.3141
P-0023-3 | 2131.1656 |  |  | 2122.1974 | 8.478 | 16.9559 | 0.0039949 | 0.3995% | 2114.3141 | 2131.1656 | 16.8515
B-0041-1 | 2136.7025
B-0041-2 | 2137.4029
B-0041-3 | 2137.7466 |  |  | 2137.284 | 0.53208 | 1.0642 | 0.00024895 | 0.0249% | 2136.7025 | 2137.7466 | 1.0441
P-0034-1 | 2112.5961
P-0034-2 | 2105.4407
P-0034-3 | 2115.5449 |  |  | 2111.1939 | 5.196 | 10.392 | 0.0024612 | 0.2461% | 2105.4407 | 2115.5449 | 10.1043
SAC-2 | 1798.3827
BAYSTD1-11172020-1 | 2167.3315
BAYSTD1-11172020-2 | 2162.0931
BAYSTD1-11172020-3 | 2163.2411 |  |  | 2164.2219 | 2.7535 | 5.5069 | 0.0012723 | 0.1272% | 2162.0931 | 2167.3315 | 5.2383
SAC-1 | 2290.3416
BAYSTD1-04072021 | 2185.6742
SAC-2 | 2450.9836
P-0084-1 | 2203.7691
P-0084-2 | 2339.1625 |  |  | 2271.4658 | 95.7376 | 191.4752 | 0.042148 | 4.215% | 2203.7691 | 2339.1625 | 135.3934
SAC-1 | 2910.6442 |  | Christian's samples
BAYSTD1-04202021-1 | 2184.3871 |  | Christian's samples
H2-702-1 | 2280.5082 |  | Christian's samples
H2-702-2 | 2251.6546 |  | Christian's samples
H2-702-3 | 2332.1848 |  | Christian's samples | 2288.1159 | 40.8005 | 81.6011 | 0.017832 | 1.783% | 2251.6546 | 2332.1848 | 80.5302
H3-702-1 | 2299.0641 |  | Christian's samples
H3-702-2 | 2306.1521 |  | Christian's samples
H3-702-3 | 2284.0452 |  | Christian's samples | 2296.4205 | 11.2881 | 22.5761 | 0.0049155 | 0.4916% | 2284.0452 | 2306.1521 | 22.1069
H2-719-1 | 2289.4798 |  | Christian's samples
H2-719-2 | 2295.8571 |  | Christian's samples | 2292.6685 | 4.5095 | 9.0189 | 0.0019669 | 0.1967% | 2289.4798 | 2295.8571 | 6.3774
SAC-1 | 2380.6236
DIC-CRM186-04272021-1 | 2250.7576 |  | not sure what happened here, might be worth excluding
DIC-CRM186-04272021-2 | 2203.2262
DIC-CRM186-04272021-3 | 2200.2796 |  | using only two values in the calulations to the right | 2201.7529 | 2.0835 | 4.1671 | 0.00094631 | 0.09463% | 2200.2796 | 2203.2262 | 2.9466
SAC-2 | 2398.0314
BAYSTD1-04272021-1 | 2188.0548
BAYSTD1-04272021-2 | 2182.1633
BAYSTD1-04272021-3 | 2179.9038 |  |  | 2183.374 | 4.2082 | 8.4164 | 0.0019274 | 0.1927% | 2179.9038 | 2188.0548 | 8.151
SAC-3 | 2110.5869
BAYSTD2-04272021-1 | 2253.4925 |  | Bay Standard Batch 2 started here
BAYSTD2-04272021-2 | 2256.3587 |  | Bay Standard Batch 2 started here
BAYSTD2-04272021-3 | 2252.4539 |  | Bay Standard Batch 2 started here | 2254.1017 | 2.0224 | 4.0448 | 0.00089721 | 0.08972% | 2252.4539 | 2256.3587 | 3.9048
SAC-1 | 2027.6002 |  | Christian's samples
BaySTD2-04282021-1 | 2251.733 |  | Christian's samples
BaySTD2-04282021-2 | 2243.0483 |  | Christian's samples | 2247.3907 | 6.141 | 12.2819 | 0.0027325 | 0.2732% | 2243.0483 | 2251.733 | 8.6846
SAC-2 | 2074.2815 |  | Christian's samples
H3-719-1 | 2293.848 |  | Christian's samples
H3-719-2 | 2299.6646 |  | Christian's samples | 2296.7563 | 4.113 | 8.2259 | 0.0017908 | 0.1791% | 2293.848 | 2299.6646 | 5.8166
T44-702-1 | 2275.8506 |  | Christian's samples
T44-702-2 | 2242.8962 |  | Christian's samples | 2259.3734 | 23.3022 | 46.6045 | 0.010314 | 1.031% | 2242.8962 | 2275.8506 | 32.9543
T4-702-1 | 2285.8484 |  | Christian's samples
T4-702-2 | 2280.6458 |  | Christian's samples | 2283.2471 | 3.6788 | 7.3577 | 0.0016112 | 0.1611% | 2280.6458 | 2285.8484 | 5.2027
T3-702-1 | 2288.1536 |  | Christian's samples
T3-702-2 | 2292.4782 |  | Christian's samples | 2290.3159 | 3.058 | 6.1159 | 0.0013352 | 0.1335% | 2288.1536 | 2292.4782 | 4.3246
SAC-1 | 2071.5652 |  | Christian's samples
BAYSTD2-05052021-1 | 2257.8633 |  | Christian's samples
BAYSTD2-05052021-2 | 2228.2726 |  | Christian's samples | 2243.0679 | 20.9237 | 41.8475 | 0.0093282 | 0.9328% | 2228.2726 | 2257.8633 | 29.5906
SAC-2 | 2009.7665 |  | Christian's samples
T6-702-1 | 2396.3344 |  | Christian's samples
T6-702-2 | 2297.7711 |  | Christian's samples | 2347.0527 | 69.6947 | 139.3895 | 0.029695 | 2.969% | 2297.7711 | 2396.3344 | 98.5632
T27-719-1 | 2309.8695 |  | Christian's samples
T27-719-2 | 2259.1674 |  | Christian's samples | 2284.5184 | 35.8518 | 71.7036 | 0.015693 | 1.569% | 2259.1674 | 2309.8695 | 50.7021
T10-719-1 | 2268.6915 |  | Christian's samples
T10-719-2 | 2265.8265 |  | Christian's samples | 2267.259 | 2.0259 | 4.0517 | 0.00089353 | 0.08935% | 2265.8265 | 2268.6915 | 2.865
T36-719-1 | 2260.527 |  | Christian's samples
T36-719-2 | 2261.1222 |  | Christian's samples | 2260.8246 | 0.42088 | 0.84176 | 0.00018616 | 0.01862% | 2260.527 | 2261.1222 | 0.59521
SAC-1 | 1989.3161 |  | Christian's samples
BaySTD2-05122021-1 | 2258.9954 |  | Christian's samples
BaySTD2-05122021-2 | 2231.564 |  | Christian's samples | 2245.2797 | 19.3969 | 38.7938 | 0.008639 | 0.8639% | 2231.564 | 2258.9954 | 27.4314
SAC-2 | 1863.6784 |  | Christian's samples
T16-719-1 | 2161.0007 |  | Christian's samples
T16-719-2 | 2267.5799 |  | Christian's samples | 2214.2903 | 75.3629 | 150.7258 | 0.034035 | 3.403% | 2161.0007 | 2267.5799 | 106.5793
T12-708-1 | 2238.043 |  | Christian's samples
T12-708-2 | 2242.358 |  | Christian's samples | 2240.2005 | 3.0512 | 6.1023 | 0.001362 | 0.1362% | 2238.043 | 2242.358 | 4.315
SAC-3 | 1652.8686 |  | Christian's samples
T6-708-1 | 2240.1269 |  | Christian's samples
T6-708-2 | 2241.2329 |  | Christian's samples | 2240.6799 | 0.78206 | 1.5641 | 0.00034903 | 0.0349% | 2240.1269 | 2241.2329 | 1.106
T20-708-1 | 2233.7405 |  | Christian's samples
T20-708-2 | 2237.4207 |  | Christian's samples | 2235.5806 | 2.6023 | 5.2046 | 0.001164 | 0.1164% | 2233.7405 | 2237.4207 | 3.6802
SAC-1 | 1688.4186 |  | Christian's samples
BaySTD2-05192021-1 | 2257.6302 |  | Christian's samples
BaySTD2-05192021-2 | 2234.1768 |  | Christian's samples | 2245.9035 | 16.584 | 33.1681 | 0.0073841 | 0.7384% | 2234.1768 | 2257.6302 | 23.4534
SAC-2 | 1586.9124 |  | Christian's samples
T8-714-1 | 2378.9742 |  | Christian's samples
T8-714-2 | 2300.7801 |  | Christian's samples | 2339.8771 | 55.2916 | 110.5833 | 0.02363 | 2.363% | 2300.7801 | 2378.9742 | 78.1942
T14-714-1 | 2284.0636 |  | Christian's samples
T14-714-2 | 2280.9784 |  | Christian's samples | 2282.521 | 2.1816 | 4.3632 | 0.00095577 | 0.09558% | 2280.9784 | 2284.0636 | 3.0852
T38-714-1 | 2271.1866 |  | Christian's samples
T38-714-2 | 2270.2599 |  | Christian's samples | 2270.7232 | 0.6553 | 1.3106 | 0.00028858 | 0.02886% | 2270.2599 | 2271.1866 | 0.92673
SAC-3 | 1464.3278 |  | Christian's samples
T21-708-1 | 2230.3629 |  | Christian's samples
T21-708-2 | 2233.7284 |  | Christian's samples | 2232.0456 | 2.3798 | 4.7596 | 0.0010662 | 0.1066% | 2230.3629 | 2233.7284 | 3.3655

## QAQC crm & baystd assessment
sample/crm | alk at 25degC | Cref value | CLabi-Cref | (xi-Cref)^2 |  | all standards |  |  |  | Dickson STD | TA | Salinity
crm130_12112019_1 | 2242.6988 | 2238.04 | 4.6588 | 21.7042 | [5 empty] | Batch 162 | 2403.72 | 33.312
crm130_12112019_2 | 2247.5364 | 2238.04 | 9.4964 | 90.181 |  | 25.0464 | RMSbias |  |  | Batch 130 | 2238.04 | 33.661
crm130_12112019_3 | 2230.1036 | 2238.04 | -7.9364 | 62.9861 |  | 0.53 | u(Cref) |  |  | Batch 186 | 2212 | 33.525
crm130_12112019_4 | 2249.3508 | 2238.04 | 11.3108 | 127.9349 |  | 25.052 | u(bias) of x | absolute uncertainty |  | BayStd1 | 2144.1721 | 27.53
crm130_12112019_5 | 2250.2916 | 2238.04 | 12.2516 | 150.1007 |  | NA | u(bias)/x | relative uncertainty
crm162_01272020_2 | 2403.7904 | 2403.72 | 0.070385 | 0.004954
crm130_01272020_1 | 2278.2466 | 2238.04 | 40.2066 | 1616.5688
crm130_01272020_2 | 2283.2413 | 2238.04 | 45.2013 | 2043.1553
crm130_01282020_1 | 2272.6035 | 2238.04 | 34.5635 | 1194.6339
crm186_01282020_2 | 2233.4604 | 2212 | 21.4604 | 460.5475
crm186_01282020_3 | 2242.3757 | 2212 | 30.3757 | 922.682
crm186_01292020_1 | 2241.6865 | 2212 | 29.6865 | 881.2912
crm186_01292020_2 | 2241.5421 | 2212 | 29.5421 | 872.7347
crm186_01292020_3 | 2236.1096 | 2212 | 24.1096 | 581.2719
crm186_01292020_4 | 2235.1851 | 2212 | 23.1851 | 537.5484
crm186_01302020_1 | 2248.0026 | 2212 | 36.0026 | 1296.1908
crm186_01202020_2 | 2247.3779 | 2212 | 35.3779 | 1251.5943
crm186_01302020_3 | 2237.5032 | 2212 | 25.5032 | 650.4133 | source: |  |  |  | source:
crm186_01202020_4 | 2238.3866 | 2212 | 26.3866 | 696.255 | https://sisu.ut.ee/sites/default/files/measurement/files/single-lab_validation_nordtest_uncertainty.pdf |  |  |  | https://www.c-can.info/reference/OAGuidanceManual.pdf
crm186_01312020_1 | 2247.372 | 2212 | 35.372 | 1251.1789
crm186_01312020_2 | 2244.8768 | 2212 | 32.8768 | 1080.8809
crm186_01312020_3 | 2241.6594 | 2212 | 29.6594 | 879.6794
crm186_01312020_4 | 2244.2021 | 2212 | 32.2021 | 1036.978
crm186_02102020_1 | 2169.8228 | 2212 | -42.1772 | 1778.9156
crm186_02102020_2 | 2189.8049 | 2212 | -22.1951 | 492.6235
crm186_02102020_3 | 2207.0936 | 2212 | -4.9064 | 24.0728
crm186_02102020_4 | 2203.9788 | 2212 | -8.0212 | 64.3392
bay001_02122020_2 | 2158.5898 | 2144.1721 | 14.4178 | 207.8719
crm186_02122020_1 | 2223.7492 | 2212 | 11.7492 | 138.0433
crm186_02122020_2 | 2228.035 | 2212 | 16.035 | 257.1205
bay001_02122020_3 | 2157.7459 | 2144.1721 | 13.5738 | 184.2491
bay001_02122020_4 | 2174.6168 | 2144.1721 | 30.4447 | 926.8818
bay001_02132020_1 | 2166.2382 | 2144.1721 | 22.0661 | 486.9117
bay001_02132020_2 | 2173.7654 | 2144.1721 | 29.5933 | 875.7631
crm186_02132020_1 | 2229.4135 | 2212 | 17.4135 | 303.2302
bay001_02132020_3 | 2192.0776 | 2144.1721 | 47.9055 | 2294.9405
bay001_02172020_1 | 2156.3511 | 2144.1721 | 12.1791 | 148.3298
bay001_02172020_2 | 2167.2566 | 2144.1721 | 23.0845 | 532.8941
bay001_02172020_3 | 2161.0007 | 2144.1721 | 16.8286 | 283.201
bay001_02172020_4 | 2173.0418 | 2144.1721 | 28.8698 | 833.4632
bay001_02172020_5 | 2182.9605 | 2144.1721 | 38.7885 | 1504.5444
bay001_02172020_6 | 2181.0465 | 2144.1721 | 36.8744 | 1359.725
crm186_02172020_1 | 2232.7276 | 2212 | 20.7276 | 429.6322
crm186_02172020_2 | 2225.3939 | 2212 | 13.3939 | 179.3979
crm186_02172020_3 | 2224.1164 | 2212 | 12.1164 | 146.8066
BAYSTD1-02192020-1 | 2172.5331 | 2144.1721 | 28.361 | 804.3481
BAYSTD1-02192020-2 | 2169.1763 | 2144.1721 | 25.0042 | 625.2092
BAYSTD1-02192020-3 | 2166.3611 | 2144.1721 | 22.189 | 492.3534
BAYSTD1-02192020-4 | 2170.5506 | 2144.1721 | 26.3785 | 695.8257
BAYSTD1-02192020-5 | 2168.9045 | 2144.1721 | 24.7324 | 611.6932
BAYSTD1-02192020-6 | 2162.9156 | 2144.1721 | 18.7435 | 351.3185
CRM186-02192020-1 | 2231.5014 | 2212 | 19.5014 | 380.3048
CRM186-02192020-2 | 2237.2213 | 2212 | 25.2213 | 636.1139
CRM186-02192020-3 | 2231.3102 | 2212 | 19.3102 | 372.8839
BAYSTD1-02202020-1 | 2171.6037 | 2144.1721 | 27.4317 | 752.496
BAYSTD1-02202020-2 | 2168.3325 | 2144.1721 | 24.1605 | 583.7278
BAYSTD1-02202020-3 | 2169.6807 | 2144.1721 | 25.5086 | 650.6882
BAYSTD1-02202020-4 | 2172.3969 | 2144.1721 | 28.2248 | 796.6387
BAYSTD1-02202020-5 | 2167.4397 | 2144.1721 | 23.2676 | 541.3805
BAYSTD1-02202020-6 | 2171.8156 | 2144.1721 | 27.6435 | 764.1638
CRM186-02202020-1 | 2234.9085 | 2212 | 22.9085 | 524.7987
CRM186-02202020-2 | 2233.8319 | 2212 | 21.8319 | 476.632
CRM186-02202020-3 | 2236.8026 | 2212 | 24.8026 | 615.1711
CRM186-11102020-1 | 2220.1115 | 2212 | 8.1115 | 65.7964
CRM186-11102020-2 | 2223.2986 | 2212 | 11.2986 | 127.659
CRM186-11102020-3 | 2217.954 | 2212 | 5.954 | 35.4495
BAYSTD1-11172020-1 | 2167.3315 | 2144.1721 | 23.1594 | 536.3576
BAYSTD1-11172020-2 | 2162.0931 | 2144.1721 | 17.9211 | 321.1642
BAYSTD1-11172020-3 | 2163.2411 | 2144.1721 | 19.069 | 363.6281

## QAQC crms assessment
sample/crm | alk at 25degC | Cref value | CLabi-Cref | (xi-Cref)^2 |  | all crms |  |  |  | Dickson STD | TA | Salinity
crm162_01272020_2 | 2403.7904 | 2403.72 | 0.070385 | 0.004954 |  | 25.933 | RMSbias |  |  | Batch 162 | 2403.72 | 33.312
crm130_12112019_1 | 2242.6988 | 2238.04 | 4.6588 | 21.7042 |  | 0.53 | u(Cref) |  |  | Batch 130 | 2238.04 | 33.661
crm130_12112019_2 | 2247.5364 | 2238.04 | 9.4964 | 90.181 |  | 25.9384 | u(bias) of x | absolute uncertainty |  | Batch 186 | 2212 | 33.525
crm130_12112019_3 | 2230.1036 | 2238.04 | -7.9364 | 62.9861 |  | NA | u(bias)/x | relative uncertainty |  | BayStd1 | 2144.1721 | 27.53
crm130_12112019_4 | 2249.3508 | 2238.04 | 11.3108 | 127.9349
crm130_12112019_5 | 2250.2916 | 2238.04 | 12.2516 | 150.1007 |  | crm130
crm130_01272020_1 | 2278.2466 | 2238.04 | 40.2066 | 1616.5688
crm130_01272020_2 | 2283.2413 | 2238.04 | 45.2013 | 2043.1553 |  | 25.7567 | RMSbias
crm130_01282020_1 | 2272.6035 | 2238.04 | 34.5635 | 1194.6339 |  | 0.53 | u(Cref)
crm186_01282020_2 | 2233.4604 | 2212 | 21.4604 | 460.5475 |  | 25.7622 | u(bias) of x | absolute uncertainty
crm186_01282020_3 | 2242.3757 | 2212 | 30.3757 | 922.682 |  | 0.011511 | u(bias)/x | relative uncertainty
crm186_01292020_1 | 2241.6865 | 2212 | 29.6865 | 881.2912
crm186_01292020_2 | 2241.5421 | 2212 | 29.5421 | 872.7347
crm186_01292020_3 | 2236.1096 | 2212 | 24.1096 | 581.2719 |  | crm162
crm186_01292020_4 | 2235.1851 | 2212 | 23.1851 | 537.5484
crm186_01302020_1 | 2248.0026 | 2212 | 36.0026 | 1296.1908 |  | 0.070385 | RMSbias
crm186_01202020_2 | 2247.3779 | 2212 | 35.3779 | 1251.5943 |  | 0.53 | u(Cref)
crm186_01302020_3 | 2237.5032 | 2212 | 25.5032 | 650.4133 | source: | 0.53465 | u(bias) of x | absolute uncertainty | source:
crm186_01202020_4 | 2238.3866 | 2212 | 26.3866 | 696.255 | https://sisu.ut.ee/sites/default/files/measurement/files/single-lab_validation_nordtest_uncertainty.pdf | 0.00022243 | u(bias)/x | relative uncertainty | https://www.c-can.info/reference/OAGuidanceManual.pdf
crm186_01312020_1 | 2247.372 | 2212 | 35.372 | 1251.1789
crm186_01312020_2 | 2244.8768 | 2212 | 32.8768 | 1080.8809 |  | crm186
crm186_01312020_3 | 2241.6594 | 2212 | 29.6594 | 879.6794
crm186_01312020_4 | 2244.2021 | 2212 | 32.2021 | 1036.978 |  | 24.2763 | RMSbias
crm186_02102020_1 | 2169.8228 | 2212 | -42.1772 | 1778.9156 |  | 0.53 | u(Cref)
crm186_02102020_2 | 2189.8049 | 2212 | -22.1951 | 492.6235 |  | 24.2821 | u(bias) of x | absolute uncertainty
crm186_02102020_3 | 2207.0936 | 2212 | -4.9064 | 24.0728 |  | 0.010977 | u(bias)/x | relative uncertainty
crm186_02102020_4 | 2203.9788 | 2212 | -8.0212 | 64.3392
crm186_02122020_1 | 2223.7492 | 2212 | 11.7492 | 138.0433
crm186_02122020_2 | 2228.035 | 2212 | 16.035 | 257.1205
crm186_02132020_1 | 2229.4135 | 2212 | 17.4135 | 303.2302
crm186_02172020_1 | 2232.7276 | 2212 | 20.7276 | 429.6322 |  | crm186 | probe C started on 11/10/2020
crm186_02172020_2 | 2225.3939 | 2212 | 13.3939 | 179.3979
crm186_02172020_3 | 2224.1164 | 2212 | 12.1164 | 146.8066 |  | 8.7351 | RMSbias
CRM186-02192020-1 | 2231.5014 | 2212 | 19.5014 | 380.3048 |  | 0.53 | u(Cref)
CRM186-02192020-2 | 2237.2213 | 2212 | 25.2213 | 636.1139 |  | 8.7511 | u(bias) of x | absolute uncertainty
CRM186-02192020-3 | 2231.3102 | 2212 | 19.3102 | 372.8839 |  | 0.0039562 | u(bias)/x | relative uncertainty
CRM186-02202020-1 | 2234.9085 | 2212 | 22.9085 | 524.7987
CRM186-02202020-2 | 2233.8319 | 2212 | 21.8319 | 476.632
CRM186-02202020-3 | 2236.8026 | 2212 | 24.8026 | 615.1711
CRM186-11102020-1 | 2220.1115 | 2212 | 8.1115 | 65.7964
CRM186-11102020-2 | 2223.2986 | 2212 | 11.2986 | 127.659
CRM186-11102020-3 | 2217.954 | 2212 | 5.954 | 35.4495

## QAQC baystds assessment
sample/crm | alk at 25degC | Cref value | CLabi-Cref | (xi-Cref)^2 | [5 empty] | Dickson STD | TA | Salinity
bay001_02122020_2 | 2158.5898 | 2144.1721 | 14.4178 | 207.8719 |  | 26.1971 | RMSbias |  |  | Batch 162 | 2403.72 | 33.312
bay001_02122020_3 | 2157.7459 | 2144.1721 | 13.5738 | 184.2491 |  | 0.53 | u(Cref) |  |  | Batch 130 | 2238.04 | 33.661
bay001_02122020_4 | 2174.6168 | 2144.1721 | 30.4447 | 926.8818 |  | 26.2025 | u(bias) of x | absolute uncertainty |  | Batch 186 | 2212 | 33.525
bay001_02132020_1 | 2166.2382 | 2144.1721 | 22.0661 | 486.9117 |  | 0.01222 | u(bias)/x | relative uncertainty |  | BayStd1 | 2144.1721 | 27.53
bay001_02132020_2 | 2173.7654 | 2144.1721 | 29.5933 | 875.7631 | [5 empty] | BayStd2 | 2254.1017 | 30.34
bay001_02132020_3 | 2192.0776 | 2144.1721 | 47.9055 | 2294.9405
bay001_02172020_1 | 2156.3511 | 2144.1721 | 12.1791 | 148.3298
bay001_02172020_2 | 2167.2566 | 2144.1721 | 23.0845 | 532.8941
bay001_02172020_3 | 2161.0007 | 2144.1721 | 16.8286 | 283.201
bay001_02172020_4 | 2173.0418 | 2144.1721 | 28.8698 | 833.4632
bay001_02172020_5 | 2182.9605 | 2144.1721 | 38.7885 | 1504.5444
bay001_02172020_6 | 2181.0465 | 2144.1721 | 36.8744 | 1359.725
BAYSTD1-02192020-1 | 2172.5331 | 2144.1721 | 28.361 | 804.3481
BAYSTD1-02192020-2 | 2169.1763 | 2144.1721 | 25.0042 | 625.2092
BAYSTD1-02192020-3 | 2166.3611 | 2144.1721 | 22.189 | 492.3534
BAYSTD1-02192020-4 | 2170.5506 | 2144.1721 | 26.3785 | 695.8257
BAYSTD1-02192020-5 | 2168.9045 | 2144.1721 | 24.7324 | 611.6932
BAYSTD1-02192020-6 | 2162.9156 | 2144.1721 | 18.7435 | 351.3185 | source: |  |  |  | source:
BAYSTD1-02202020-1 | 2171.6037 | 2144.1721 | 27.4317 | 752.496 | https://sisu.ut.ee/sites/default/files/measurement/files/single-lab_validation_nordtest_uncertainty.pdf |  |  |  | https://www.c-can.info/reference/OAGuidanceManual.pdf
BAYSTD1-02202020-2 | 2168.3325 | 2144.1721 | 24.1605 | 583.7278
BAYSTD1-02202020-3 | 2169.6807 | 2144.1721 | 25.5086 | 650.6882
BAYSTD1-02202020-4 | 2172.3969 | 2144.1721 | 28.2248 | 796.6387
BAYSTD1-02202020-5 | 2167.4397 | 2144.1721 | 23.2676 | 541.3805 |  | baystd1 | probe C started on 11/10/2020
BAYSTD1-02202020-6 | 2171.8156 | 2144.1721 | 27.6435 | 764.1638
BAYSTD1-11172020-1 | 2167.3315 | 2144.1721 | 23.1594 | 536.3576 |  | 20.1755 | RMSbias
BAYSTD1-11172020-2 | 2162.0931 | 2144.1721 | 17.9211 | 321.1642 |  | 0.53 | u(Cref)
BAYSTD1-11172020-3 | 2163.2411 | 2144.1721 | 19.069 | 363.6281 |  | 20.1824 | u(bias) of x | absolute uncertainty
BAYSTD1-04072021 | 2185.6742 | 2144.1721 | 41.5022 | 1722.4297 |  | 0.0094127 | u(bias)/x | relative uncertainty

## Bay Std second look
sample/crm | alk at 25degC | Cref value | CLabi-Cref | (xi-Cref)^2 |  |  |  | Dickson STD | TA | Salinity
bay001_02122020_2 | 2158.5898 | 2163.6509 | -5.061 | 25.6139 | 9.4721 | RMSbias |  | Batch 162 | 2403.72 | 33.312
bay001_02122020_3 | 2157.7459 | 2163.6509 | -5.9049 | 34.8683 | 0.53 | u(Cref) |  | Batch 130 | 2238.04 | 33.661
bay001_02122020_4 | 2174.6168 | 2163.6509 | 10.966 | 120.2522 | 9.487 | u(bias) of x | absolute uncertainty | Batch 186 | 2212 | 33.525
bay001_02132020_1 | 2166.2382 | 2163.6509 | 2.5873 | 6.6941 | 0.0043847 | u(bias)/x | relative uncertainty | BayStd1 | 2163.6509 | 27.53
bay001_02132020_2 | 2173.7654 | 2163.6509 | 10.1145 | 102.3034
bay001_02132020_3 | 2192.0776 | 2163.6509 | 28.4268 | 808.0806
bay001_02172020_1 | 2156.3511 | 2163.6509 | -7.2997 | 53.2857
bay001_02172020_2 | 2167.2566 | 2163.6509 | 3.6057 | 13.0012
bay001_02172020_3 | 2161.0007 | 2163.6509 | -2.6502 | 7.0236
bay001_02172020_4 | 2173.0418 | 2163.6509 | 9.391 | 88.1906
bay001_02172020_5 | 2182.9605 | 2163.6509 | 19.3097 | 372.8637
bay001_02172020_6 | 2181.0465 | 2163.6509 | 17.3957 | 302.6094
BAYSTD1-02192020-1 | 2172.5331 | 2163.6509 | 8.8823 | 78.8944
BAYSTD1-02192020-2 | 2169.1763 | 2163.6509 | 5.5254 | 30.5301
BAYSTD1-02192020-3 | 2166.3611 | 2163.6509 | 2.7103 | 7.3455
BAYSTD1-02192020-4 | 2170.5506 | 2163.6509 | 6.8997 | 47.6063
BAYSTD1-02192020-5 | 2168.9045 | 2163.6509 | 5.2537 | 27.6009
BAYSTD1-02192020-6 | 2162.9156 | 2163.6509 | -0.73529 | 0.54064
BAYSTD1-02202020-1 | 2171.6037 | 2163.6509 | 7.9529 | 63.2483
BAYSTD1-02202020-2 | 2168.3325 | 2163.6509 | 4.6817 | 21.9181
BAYSTD1-02202020-3 | 2169.6807 | 2163.6509 | 6.0298 | 36.3586
BAYSTD1-02202020-4 | 2172.3969 | 2163.6509 | 8.746 | 76.4927
BAYSTD1-02202020-5 | 2167.4397 | 2163.6509 | 3.7888 | 14.3551 | baystd1 | probe C started on 11/10/2020
BAYSTD1-02202020-6 | 2171.8156 | 2163.6509 | 8.1647 | 66.6629
BAYSTD1-11172020-1 | 2167.3315 | 2163.6509 | 3.6806 | 13.5469 | 2.3196 | RMSbias
BAYSTD1-11172020-2 | 2162.0931 | 2163.6509 | -1.5577 | 2.4265 | 0.53 | u(Cref)
BAYSTD1-11172020-3 | 2163.2411 | 2163.6509 | -0.40974 | 0.16789 | 2.3794 | u(bias) of x | absolute uncertainty
[5 empty] | 0.0010997 | u(bias)/x | relative uncertainty