## pH to mV from calibration
Nernst Slope at 25degC with offsets
pH | mV from calibration | "+25mV" | "-25mV"
4 | 177.48 | 202.48 | 152.48 |  | slope1 | -59.16 | slope1% | 100 |  | acceptable probe offset from 0mV at pH7
7 | 0 | 25 | -25 |  | slope2 | -59.16 | slope2% | 100 |  | "+- 25 mV"
10 | -177.48 | -152.48 | -202.48
3 | 236.64 | 261.64 | 211.64 | [6 empty] | acceptable slope offset
3.5 | 207.06 | 232.06 | 182.06 | [6 empty] | 85-105%
Use calibration generated slope to calculate mV
Enter from calibration
|  | no offset at pH7 | offset at pH 7
-4.7
pH | slope | mV | mV |  | slope% |  | dispense increment Volume |  |  |  | sample id | cup # | cup mass | sample +cup | sample mass | sample #
4 | 57.05 | 171.15 | 166.45 |  | 96.4334 |  | 0.02 | mL |  |  | SAC-1 | 5 | 7.291 | 58.706 | 51.415 | 1
7 |  | 0 | -4.7 |  |  |  | 0.4 | 0.05 | 0.02 |  | BaySTD2-1 | 6 | 7.289 | 58.121 | 50.832 | 2
10 |  | -171.15 | -175.85 | [7 empty] | BaySTD2-2 | 7 | 7.279 | 57.855 | 50.576 | 3
3.5 |  | 199.675 | 194.975 | Gran titration "upper bound" aka Ehigh |  |  | first aliquot |  |  |  | SAC-2 | 8 | 7.27 | 58.659 | 51.389 | 4
3 |  | 228.2 | 223.5 | Gran titration "lower bound" aka Elow |  |  | 1.2 | mL |  |  | T8-714-1 | 9 | 7.284 | 57.566 | 50.282 | 5
3.7 |  | 188.265 | 183.565 | Enter This value to end first titration, round down to next mV | [6 empty] | T8-714-2 | 10 | 7.283 | 57.236 | 49.953 | 6
2.8 |  | 239.61 | 234.91 | Enter This value to end gran titration, round up to next mV | [6 empty] | T14-714-1 | 11 | 7.299 | 58.023 | 50.724 | 7
[11 empty] | T14-714-2 | 12 | 7.278 | 57.674 | 50.396 | 8
[11 empty] | T38-714-1 | 13 | 7.282 | 58.154 | 50.872 | 9
[11 empty] | T38-714-2 | 14 | 7.282 | 58.088 | 50.806 | 10
[11 empty] | SAC-3 | 15 | 7.301 | 58.564 | 51.263 | 11
[11 empty] | T21-708-1 | 16 | 7.302 | 57.809 | 50.507 | 12
[11 empty] | T21-708-2 | 17 | 7.303 | 58.065 | 50.762 | 13

## sample metadata
sample id | sample mass | salinity | in situ temperature | ETris | pHTris | EHigh | ELow
SAC-1 | 51.415 | 30 | 25 | -4.7 | 7 | 194.975 | 223.5
BaySTD2-1 | 50.832 | 30.34 | 14.11 | -4.7 | 7 | 194.975 | 223.5
BaySTD2-2 | 50.576 | 30.34 | 14.11 | -4.7 | 7 | 194.975 | 223.5
SAC-2 | 51.389 | 27.5 | 25 | -4.7 | 7 | 194.975 | 223.5
T8-714-1 | 50.282 | 27.06 | 15.22 | -4.7 | 7 | 194.975 | 223.5
T8-714-2 | 49.953 | 27.06 | 15.25 | -4.7 | 7 | 194.975 | 223.5
T14-714-1 | 50.724 | 28.27 | 15.71 | -4.7 | 7 | 194.975 | 223.5
T14-714-2 | 50.396 | 28.26 | 15.72 | -4.7 | 7 | 194.975 | 223.5
T38-714-1 | 50.872 | 27.06 | 15.26 | -4.7 | 7 | 194.975 | 223.5
T38-714-2 | 50.806 | 27.06 | 15.28 | -4.7 | 7 | 194.975 | 223.5
SAC-3 | 51.263 | 25 | 25 | -4.7 | 7 | 194.975 | 223.5
T21-708-1 | 50.507 | 25.12 | 18.36 | -4.7 | 7 | 194.975 | 223.5
T21-708-2 | 50.762 | 25.16 | 18.31 | -4.7 | 7 | 194.975 | 223.5